## dpSwitch
Calibration Certificate: Differential Pressure switch calibration
Calibration certificate No.: Differential Pressure Switch 1234-41 | [14 empty] | Page: 1/1
Service Request / Work Order No.: |  |  |  |  | 71234567
Description identification of instrument to be calibrated
Name: Differential Pressure Switch Calibration | [8 empty] | Tag no |  |  | PS-234-56
Sl. No.: 123456789 | [8 empty] | Make: |  |  | Danfoss
Model / Type No.: DP Switch LTX1234 - XE-11-22 | [8 empty] | Range: |  |  | 40 to 170 mbar
Sensor: Differential pressure switch | [8 empty] | Resolution: |  |  | N/A
Location: Main cooling water pressure control valve FCV-2063
Applicable specification of instrument to be calibrated: Set point, reset point and deadband within permissible limit | [10 empty] | 1%
Date of receipt of item: |  |  |  | 2026-09-30 |  |  |  | Date of calibration: |  |  |  | 2026-09-30
Calibration valid upto: |  |  |  | 2027-09-30 |  |  |  | Frequency of calibration: |  |  |  | 12 months
Environmental condition during calibration |  |  |  | Temperature: |  |  | 25°C ± 2°C
|  |  |  | Humidity: |  |  | 64% RH
Details of standard instruments Used:
Name of the instrument |  | Sr.No |  |  |  | Make/Model |  |  | Valid up to |  |  |  | Report no: / Certificate no:
Pressure calibrator |  | 1234598 |  |  |  | Fluke / 718 |  |  | 2026-09-08 |  |  |  | 1234STRW432
Digital multimeter |  | 6789123 |  |  |  | Fluke / 117 |  |  | 2026-09-08 |  |  |  | 1234MTRW432
Basis of calibration: |  |  |  | Pressure comparison with standards and verification of switch action
Applied Procedure no : XXXX122334
As found value |  |  | As left value |  |  |  |  | Expected value |  |  |  | Error value |  |  | Error in %
SET (mbar) | RESET (mbar) |  | SET (mbar) |  |  | RESET (mbar) |  | SET (mbar) |  | RESET (mbar) |  | DEADBAND (mbar) |  | RESET (mbar) | SET | RESET
50.34 | 40.65 |  | 50.01 |  |  | 40.83 |  | 50 |  | 40.8 |  | 0.01 |  | -0.03 | 0.0076923 | -0.023077
Remarks: Calibration completed successfully. Set point, reset point and deadband were verified with the low pressure side open to atmosphere. The switching values are within the acceptable limit. Maximum as left error was -0.03 mbar.
Calibrated by | [7 empty] | Checked & approved by
Name: XXXXXXXXX | [7 empty] | Name: xxxxxxxxxxxxxx
Signature: | [7 empty] | Signature:
Date: 30/9/2026 | [7 empty] | Date: 30/9/2026