## Audit Checklist Detailed
Internal Audit Checklist for ISO Process Instrument Calibration in Process Industries
Plant / Unit: |  |  |  | Audit Date: | 2026-03-16 | [6 empty] | Auditor(s):
ID | Section | Audit Area | Audit Question | Sub Checks / Detailed Points | Evidence Required | Evidence Examples | Reference Standard | Risk/Criticality | Priority | Sample Instrument Tag | Instrument Type | Location | Last Cal Date | Next Due Date | Audit Frequency | Compliance (Yes/No) | Finding (Y/N) | Finding Details | Severity | Corrective Action | Responsible | Target Completion Date | Verification Date | Auditor Initials | Status | Notes
AC-001 | Calibration Management System | Instrument Master List | Is there an up to date instrument master list with unique tag numbers? | Master list contains tag, location, instrument type, calibration interval and owner | Master instrument database export | Master list CSV/printout | ISO 9001 | Critical | P1 | P-101 | Pressure Transmitter | Compressor House | 2025-08-28 | 2026-08-28 | Annual
AC-002 | Calibration Management System | Instrument Master Linkage | Is each instrument linked to its calibration record in CMMS? | Sample cross reference for random tags | Calibration records linked to tag in CMMS | CMMS work order and certificate PDF | ISO 9001 | High | P2 | FT-201 | Coriolis Flow Meter | Reactor Feed | 2025-02-09 | 2026-02-09 | Annual
AC-003 | Calibration Management System | Criticality Classification | Are instruments classified by criticality (safety, custody, quality)? | Criticality matrix, documented criteria | Criticality register and examples | Criticality matrix document | ISO 10012 | Critical | P1 | P-102 | Temp Transmitter | Furnace Zone | 2025-12-06 | 2026-12-06 | Annual
AC-004 | Calibration Management System | Policy & Roles | Is there a documented calibration policy and defined responsibilities? | Procedure with roles and authority | Signed procedure and org chart | Quality manual / calibration policy | ISO 9001 | Medium | P3 | [5 empty] | Annual
AC-005 | Calibration Management System | Schedule Monitoring | Is the calibration schedule implemented and monitored for overdue items? | Schedule reports, overdue lists | Schedule reports, overdue lists | CMMS overdue report export | ISO 9001 | High | P2 | FT-201 | Coriolis Flow Meter | Reactor Feed | 2025-02-09 | 2026-02-09 | Monthly
AC-006 | Calibration Procedures | Approved Procedures | Are approved calibration procedures available for each instrument type? | Controlled procedures with revision history | Procedure documents with approval signatures | Procedure document pages showing approval | ISO 10012 | High | P2 | [5 empty] | As Required
AC-007 | Calibration Procedures | Procedure Content | Do procedures include acceptance tolerances, test points and as found/as left format? | Procedure template and example reports | Procedure template and example reports | Procedure excerpt, sample report | Plant Procedure | High | P2 | [5 empty] | As Required
AC-008 | Calibration Procedures | Procedure Review | Are procedures periodically reviewed and version controlled? | Revision history and review dates | Revision log and change records | Revision log | Document Control | Medium | P3 | [5 empty] | Annual
AC-009 | Calibration Procedures | Field Method Validation | Are field calibration methods validated or verified? | Method validation reports or comparison tests | Validation records or comparison tests | Comparison against lab results | ISO 17025 | Medium | P3 | [5 empty] | As Needed
AC-010 | Calibration Equipment | Reference Standard Register | Are reference standards listed with IDs and due dates? | Reference standard register with IDs and next cal date | Register and certificates | Register export, sticker IDs | ISO 17025 | Critical | P1 | REF-PT-01 | Pressure Calibrator | Calibration Lab | 2024-07-24 | 2026-05-15 | Annual
AC-011 | Calibration Equipment | Reference Standards Validity | Are all reference standards within calibration due date? | Certificates and expiry dates | Certificates with next due date | Certificate PDF | ISO 17025 | Critical | P1 | REF-FLOW-01 | Master Flow Calibrator | Calibration Lab | 2025-03-31 | 2026-03-31 | Annual
AC-012 | Calibration Equipment | Calibration Lab Competence | Are reference standards calibrated by competent/accredited labs or traceable in house? | Lab accreditation or internal traceability chain | Accreditation or traceability docs | Lab accreditation certificate | ISO/IEC 17025 | Critical | P1 | [5 empty] | Annual
AC-013 | Calibration Equipment | Storage & Protection | Are calibration rigs and portable calibrators stored correctly? | Storage procedures and photos | Storage procedure, photos | Photos of cabinet, environment controls | Plant SOP | Low | P4 | [5 empty] | Ongoing
AC-014 | Calibration Execution | As Found / As Left Recording | Are technicians recording as found and as left values with test conditions? | Calibration reports with as found/as left and environmental conditions | Sample calibration reports | Sample certificate pages | Plant Procedure | High | P2 | P-101 | Pressure Transmitter | Compressor House | 2025-08-28 | 2026-08-28 | Per Calibration
AC-015 | Calibration Execution | Representative Test Points | Are test points representative of normal plant operating range? | List of test points vs process range | Comparison table in report | Report excerpt | Plant Procedure | High | P2 | [5 empty] | Per Calibration
AC-016 | Calibration Execution | Zero and Span Checks | Are zero and span checks performed and documented? | Reports showing zero/span adjustments | Calibration reports showing zero/span | Calibration certificate entries | Plant Procedure | Medium | P3 | [5 empty] | Per Calibration
AC-017 | Calibration Execution | Control Valve Checks | Are control valves and positioners function checked and stroked as part of calibration? | Valve test reports and stroke charts | Valve test reports | Positioner bench tests and stroke logs | Plant Procedure | High | P2 |  | Control Valve | Various | 2025-09-17 | 2026-09-17 | Annual
AC-018 | Calibration Execution | Supervisor Witnessing | Are field calibrations witnessed or spot checked by supervisor? | Witness sign off and spot check logs | Supervisor checklist signature | Supervisor sign off | Plant Procedure | Medium | P3 | [5 empty] | Periodic
AC-019 | Documentation | Certificate Contents | Do calibration certificates include uncertainty and reference standard IDs? | Sample certificates showing uncertainty and ref IDs | Sample certificate with uncertainty | PDF certificate pages | ISO 17025 / ISO 9001 | Critical | P1 | FT-201 | Coriolis Flow Meter | Reactor Feed | 2025-02-09 | 2026-02-09 | Annual
AC-020 | Documentation | Certificate Archival | Are calibration certificates archived in DMS? | DMS index entries and scanned files | DMS screenshot | DMS screenshot showing certificate | ISO 9001 | High | P2 | [5 empty] | Ongoing
AC-021 | Documentation | Out of Service Records | Is there a record of instruments removed from service and reason? | Out of service log with tag, date and reason | Out of service work order in CMMS | Out of service record | Plant SOP | Medium | P3 | [5 empty] | As Occurs
AC-022 | Traceability | Traceability Chain | Is there an unbroken chain of traceability to national metrology institute? | Traceability chain documents and certs | Traceability documents | Calibration lab traceability statement | ISO 17025 | Critical | P1 | [5 empty] | Annual
AC-023 | Traceability | Uncertainty Budgets | Are uncertainty budgets prepared for critical measurements? | Uncertainty calculations and reports | Uncertainty calculation docs | Uncertainty spreadsheet for sample instrument | ISO 14253 | High | P2 | [5 empty] | Annual
AC-024 | Labeling & Identification | Calibration Labels | Are calibration labels/stickers applied with dates? | Photographs or inspection checklist | Photo with visible sticker | Photo evidence | Plant Inspection | Low | P4 | [5 empty] | Per Calibration
AC-025 | Labeling & Identification | Tag Legibility | Is physical tag number legible and matching master list? | Field photos and master list cross check | Photo and master list row | Inspection photo | Plant Inspection | Low | P4 | [5 empty] | Periodic
AC-026 | Non Conformance Handling | Out of Tolerance Process | Is there a process for identifying out of tolerance instruments? | NCR records and investigation reports | NCR form and investigation notes | NCR evidence | Quality Procedure | Critical | P1 | [5 empty] | As Occurs
AC-027 | Non Conformance Handling | Impact Review | Are measurement data produced while instrument was out of tolerance reviewed? | Data impact assessment and product disposition records | Data review report and batch disposition | Data review and disposition | Quality / Production Records | Critical | P1 | [5 empty] | As Occurs
AC-028 | Non Conformance Handling | CAPA Tracking | Are corrective actions assigned, tracked and verified? | CAPA log with owners and verification evidence | CAPA entries in quality system | CAPA screenshots | Quality Procedure | High | P2 | [5 empty] | Ongoing
AC-029 | Software / Digital Systems | CMMS Usage | Is CMMS used to schedule and record calibrations? | CMMS screenshots of calibration work orders | CMMS export and reports | CMMS work order screenshots | ISO 9001 | High | P2 | [5 empty] | Ongoing
AC-030 | Software / Digital Systems | Access Control | Is access to calibration records protected and audited? | User access matrix and audit logs | System audit log export | Audit log screenshot | IT / Quality | High | P2 | [5 empty] | Annual
AC-031 | Software / Digital Systems | Backups | Are backups performed for calibration DB and DMS? | Backup schedule and last backup proof | Backup logs | IT backup evidence | IT Policy | Medium | P3 | [5 empty] | Weekly
AC-032 | Technician Competency | Training Records | Are technicians trained on instruments and procedures? | Training matrix and certificates | Training completion certificates | Training certificates | HR / Training Records | High | P2 | [5 empty] | Annual
AC-033 | Technician Competency | Practical Assessment | Is there evidence of practical competency assessment? | Competency assessment forms and witnessed calibrations | Signed competency form | Competency assessments | HR / Quality | Medium | P3 | [5 empty] | Periodic
AC-034 | Data Integrity & Review | Supervisor Review | Are calibration results reviewed and signed by competent verifier? | Review signatures on certificates or electronic approval | Signed certificate pages | Approval signature | Quality Procedure | High | P2 | [5 empty] | Per Calibration
AC-035 | Data Integrity & Review | Electronic Audit Trails | Are electronic calibration data protected against alteration? | Audit trails and file properties | System audit log extract | Audit log showing changes | IT / CMMS | High | P2 | [5 empty] | Annual
AC-036 | Physical Storage & Handling | Controlled Storage | Are spares and standards stored in controlled environment? | Storage condition records and photos | Photo of storage area and temp log | Storage images and logs | Maintenance / Calibration Lab | Medium | P3 | [5 empty] | Ongoing
AC-037 | Physical Storage & Handling | Transport Protection | Are transport procedures in place to protect standards during movement? | Transport SOP and logs | Transport checklists | Transport logs | Logistics / Calibration Lab | Low | P4 | [5 empty] | As Occurs
AC-038 | Vendor Management | Vendor Approval | Are external calibration vendors approved and evaluated? | Supplier evaluation and accreditation evidence | Approved vendor list and audit report | Vendor accreditation docs | Procurement | High | P2 | [5 empty] | Annual
AC-039 | Vendor Management | Certificate Verification | Are vendor certificates checked for completeness and traceability? | Incoming certificate check sheet and QC sign off | Certificate review form | Certificate review form | Quality / Calibration Lab | High | P2 | [5 empty] | Per Delivery
AC-040 | Custody Transfer & Financial Measurement | Custody Policy | Are custody transfer meters on stricter calibration schedules? | Custody transfer policy and frequency evidence | Policy document and schedule | Custody policy doc | Commercial / Operations | Critical | P1 | [5 empty] | Quarterly
AC-041 | Custody Transfer & Financial Measurement | Proving Records | Are meter proving and validation records available and current? | Proving reports and master meter records | Proving certificates | Proving reports | Operations / Flow Lab | Critical | P1 | [5 empty] | Per Proving
AC-042 | Analyzers & Continuous Emission | Analyzer Span/Zero | Are analyzers calibrated with appropriate span gases and zero checks? | Analyzer calibration logs and COA for gases | Analyzer reports and gas COA | Analyzer calibration log | Environmental / Lab | Critical | P1 | Gas-AN-01 | Gas Analyzer | Stack | 2026-01-15 | 2026-04-15 | Monthly
AC-043 | Analyzers & Continuous Emission | Sampling System Check | Are sample conditioning and lines checked during calibration? | Sampling system inspection checklist | Sampling line inspection photos | Inspection checklist | Environmental Procedure | High | P2 | [5 empty] | Per Calibration
AC-044 | Sampling & Environmental Conditions | Record Test Conditions | Are calibration test conditions recorded (temp, pressure, humidity)? | Calibration report environmental fields | Report showing environmental conditions | Report screenshot | Plant Procedure | Medium | P3 | [5 empty] | Per Calibration
AC-045 | Sampling & Environmental Conditions | Field Checks Plan | Is there a plan to perform field checks under representative conditions? | Field check plan and executed records | Field check reports | Field check report | Maintenance / QA | Medium | P3 | [5 empty] | Periodic
AC-046 | Safety & Lockout | LOTO Procedures | Are lockout and isolation procedures used during calibrations requiring line breaks? | LOTO permits and isolation certificates | Signed LOTO permit | LOTO permit photo | Safety / Maintenance | Critical | P1 | [5 empty] | Per Job
AC-047 | Safety & Lockout | Safety Prechecks | Are safety checks completed before calibration (pressure relief, purge)? | Safety checklists and sign off | Safety checklist with signatures | Signed safety checklist | Safety Procedure | Critical | P1 | [5 empty] | Per Job
AC-048 | Spare Parts & Rotables | Spare Tracking | Are spare instruments tracked and calibration status known before commissioning? | Spare inventory with calibration status | Spare parts register entry | Spare register | Stores / Maintenance | Medium | P3 | [5 empty] | Ongoing
AC-049 | Continuous Improvement | Trend Analysis | Are calibration trends analyzed to detect drift or systemic issues? | Trend charts, control charts or statistical analysis | Trend report and charts | Trend plots | Reliability / QA | Medium | P3 | [5 empty] | Quarterly
AC-050 | Continuous Improvement | Lessons Learned | Are lessons from audits incorporated into procedures and training? | Procedure updates and training logs after audit | Revision log and training records | Training records post audit | Quality / Training | Low | P4 | [5 empty] | After Audit
AC-051 | Audit Specific | Previous Audit Closure | Were previous audit non conformities closed and verified? | Previous audit report and closure evidence | Closure records and verification | Audit closure evidence | Quality | High | P2 | [5 empty] | Per Audit
AC-052 | Audit Specific | Spot Checks | Are random spot checks performed by internal audit? | Spot check log and results | Spot check sample records | Spot check logs | Internal Audit | Medium | P3 | [5 empty] | Periodic
AC-053 | Records Completeness | Certificate Fields | Do certificates contain ID, date, results, uncertainty, ref standard, and technician? | Sample certificate review checklist | Reviewed certificate showing required fields | Certificate pages | ISO 17025 / ISO 9001 | Critical | P1 | [5 empty] | Per Certificate
AC-054 | Legal & Regulatory | Regulatory Instruments | Are regulatory reporting instruments on a controlled calibration program? | Regulatory instrument register and calibration history | Register and certificate evidence | Register entries | Environmental Regulation | Critical | P1 | [5 empty] | Quarterly
AC-055 | Business Continuity | Contingency Plan | Is there a contingency plan for calibration during shutdowns or vendor unavailability? | Contingency plans and records of execution | Plan document and executed instances | Contingency record | Operations / Maintenance | Medium | P3 | [5 empty] | As Required
AC-056 | Audit Closure | CAPA Capture | Are corrective actions from current audit captured with owner and verification date? | CAPA entries with verification evidence | CAPA log entries | CAPA system screenshot | Quality | High | P2 | [5 empty] | Immediate
Instructions: Use filters to focus on critical or overdue items. Populate compliance and findings during audit walkdowns. Track CAPA in Corrective Action fields.