## Checklist
Analyzer and Sampling System Running Inspection Checklist
Detailed process-industry checklist for analyzer power, sample flow, filters, condensation, calibration, alarms, and diagnostics
Item ID | Section | Inspection Point | Detailed Check | Acceptance Criteria / Action | Status | Priority | Owner | Remarks
Safety and Pre-Checks
AI-01 | Safety and Pre-Checks | Verify work permit and job authorization | Confirm permit requirement, lockout/tagout, and plant work approval before approaching the analyzer system. | Permits valid and site access approved | In Progress | High
AI-02 | Safety and Pre-Checks | Check PPE and hazardous area requirements | Wear gloves, goggles, helmet, flame-resistant clothing, and use intrinsically safe tools where required. | Correct PPE and tools available |  | High
AI-03 | Safety and Pre-Checks | Verify analyzer shelter ventilation | Confirm shelter ventilation fan, air exchange, and toxic gas protection systems are operating correctly. | Ventilation healthy and no gas buildup |  | Critical
AI-04 | Safety and Pre-Checks | Check calibration gas cylinder safety | Ensure cylinders are secured, labeled, in date, and connected safely with no leak risk. | Cylinders secure and certification valid |  | Critical
Analyzer Power and Health
AI-05 | Analyzer Power and Health | Verify main power supply | Check incoming power, breaker status, supply healthy indication, and backup power if applicable. | Analyzer powered normally |  | High
AI-06 | Analyzer Power and Health | Inspect display and status indications | Confirm display is readable, active, and free from abnormal fault or startup messages. | No abnormal errors on display |  | High
AI-07 | Analyzer Power and Health | Review diagnostic messages | Check internal alarms, warning codes, and health messages for early problem detection. | No unresolved diagnostic alarms |  | High
AI-08 | Analyzer Power and Health | Confirm DCS or PLC communication | Verify signal communication, data refresh, and alarm transmission to control system. | Communication stable |  | Critical
AI-09 | Analyzer Power and Health | Confirm warm-up completion | Make sure the analyzer has completed start-up stabilization before trusting readings. | Analyzer fully warmed up |  | High
Sample Flow Verification
AI-10 | Sample Flow Verification | Check rotameter or flow indicator | Verify sample flow is within the recommended operating range shown by the instrument. | Flow in normal range |  | High
AI-11 | Sample Flow Verification | Look for low flow condition | Check whether the flow is below normal and investigate potential restriction or blockage. | No low flow condition |  | Critical
AI-12 | Sample Flow Verification | Check for excessive flow | Confirm flow is not high enough to overload the sensor or disturb measurement stability. | No excessive flow |  | High
AI-13 | Sample Flow Verification | Inspect bypass or vent flow | Verify bypass, return, or vent lines are operating correctly and not restricted. | Bypass and vent paths clear |  | Medium
AI-14 | Sample Flow Verification | Look for pulsating flow | Confirm sample flow is steady without oscillation caused by pressure or regulator issues. | Stable flow observed |  | Medium
Filters and Regulators
AI-15 | Filters and Regulators | Inspect particulate filters | Check for dirt loading, discoloration, moisture, and abnormal pressure drop across filters. | Filters clean or replaced when needed |  | Critical
AI-16 | Filters and Regulators | Inspect coalescing filters | Verify liquid separation is effective and there is no carryover or saturation. | No liquid carryover |  | Medium
AI-17 | Filters and Regulators | Inspect pressure regulators | Check outlet pressure stability and confirm the regulator does not hunt or fluctuate. | Stable outlet pressure |  | High
AI-18 | Filters and Regulators | Check differential pressure indicators | Review indicators across filters to identify clogging before analyzer performance is affected. | Pressure drop within limit |  | Medium
AI-19 | Filters and Regulators | Verify moisture separators and drains | Make sure drains are open and condensate is removed without blockage. | Drain system functioning |  | Medium
Condensation and Moisture Control
AI-20 | Condensation and Moisture Control | Inspect sample lines for visible condensation | Look for water droplets, cloudy tubing, or liquid accumulation in any part of the line. | No visible condensation |  | Critical
AI-21 | Condensation and Moisture Control | Verify heat tracing operation | Confirm heaters, tracing power, and temperature maintenance are functioning properly. | Heat tracing healthy |  | High
AI-22 | Condensation and Moisture Control | Check insulation and heated enclosures | Inspect insulation condition and shelter heating for cold spots or heat loss. | Sample temperature maintained |  | Medium
AI-23 | Condensation and Moisture Control | Inspect moisture traps and condensate pots | Drain and verify moisture traps are not full or plugged. | Moisture removed effectively |  | Medium
AI-24 | Condensation and Moisture Control | Check sample cooler performance | Verify cooler operation where installed to prevent unwanted vapor condensation downstream. | Cooler working normally |  | Medium
Analyzer Reading Stability
AI-25 | Analyzer Reading Stability | Observe reading stability | Watch analyzer output for abnormal fluctuation, drift, or sudden spikes during normal operation. | Stable reading trend |  | High
AI-26 | Analyzer Reading Stability | Compare with process conditions | Check whether analyzer behavior matches expected process changes and operating state. | Reading consistent with process |  | Medium
AI-27 | Analyzer Reading Stability | Review DCS trending | Use trend analysis to spot slow drift, flat signal, lagging response, or intermittent behavior. | Trend normal |  | Medium
AI-28 | Analyzer Reading Stability | Compare with laboratory result if available | Validate online reading against lab sample or offline reference value. | Deviation within acceptable limit |  | Medium
AI-29 | Analyzer Reading Stability | Confirm analyzer response time | Check whether the analyzer responds promptly to process changes without excessive delay. | Response time acceptable |  | Medium
Tubing, Fittings, and Leak Checks
AI-30 | Tubing, Fittings, and Leak Checks | Inspect tube fittings | Check compression fittings, joints, and seals for looseness or damage. | No loose fittings |  | High
AI-31 | Tubing, Fittings, and Leak Checks | Inspect flexible hoses and valve connections | Confirm hoses, valves, and connections are not cracked, brittle, or distorted. | Connections sound |  | Medium
AI-32 | Tubing, Fittings, and Leak Checks | Check tubing supports | Verify tubing is routed correctly, protected from vibration, and properly supported. | Tubing well supported |  | Medium
AI-33 | Tubing, Fittings, and Leak Checks | Inspect corrosion and mechanical damage | Look for corrosion, crushing, bending, or wear on lines and hardware. | No damage present |  | Medium
AI-34 | Tubing, Fittings, and Leak Checks | Perform leak check | Use approved leak detection methods to confirm the sample path is leak free. | No leakage detected |  | Critical
Calibration Verification
AI-35 | Calibration Verification | Check calibration gas availability | Confirm zero and span gases are available, connected, and within valid expiry. | Calibration gas available |  | Medium
AI-36 | Calibration Verification | Verify cylinder pressure | Ensure cylinder pressure is sufficient for planned calibration or verification. | Cylinder pressure adequate |  | Medium
AI-37 | Calibration Verification | Check calibration due date | Confirm calibration schedule is current and not overdue. | Calibration in date |  | High
AI-38 | Calibration Verification | Review auto calibration sequence | Verify automatic calibration starts, completes, and returns to online service correctly. | Auto calibration healthy |  | Medium
AI-39 | Calibration Verification | Check analyzer drift | Review drift between calibrations and identify sensor aging or contamination. | Drift within acceptable range |  | Medium
AI-40 | Calibration Verification | Verify calibration records | Ensure logs are updated and traceable for maintenance and quality requirements. | Records updated |  | Medium
Alarms, Interlocks, and Communications
AI-41 | Alarms, Interlocks, and Communications | Check high and low alarms | Confirm alarm thresholds, annunciation, and operator response work correctly. | Alarm functions verified |  | Critical
AI-42 | Alarms, Interlocks, and Communications | Check sample flow alarm | Verify low flow alarm activates when the sample stream drops below acceptable limits. | Flow alarm operational |  | Critical
AI-43 | Alarms, Interlocks, and Communications | Verify shelter gas detection alarms | Check gas detection alarms and any shutdown or warning actions if applicable. | Shelter alarms functional |  | Critical
AI-44 | Alarms, Interlocks, and Communications | Inspect interlock actions | Confirm interlock logic behaves correctly and protects the analyzer and plant. | Interlocks respond correctly |  | Medium
AI-45 | Alarms, Interlocks, and Communications | Verify signal transmission to DCS or PLC | Check measurement, alarm, and fault signals are reaching the control system. | Signal transmission stable |  | Medium
Advanced Diagnostics and Predictive Monitoring
AI-46 | Advanced Diagnostics and Predictive Monitoring | Review internal analyzer diagnostics | Check sensor health, internal temperature, flow alarms, and calibration deviation tracking. | Diagnostics healthy |  | Medium
AI-47 | Advanced Diagnostics and Predictive Monitoring | Inspect differential pressure trends | Use pressure drop trends across filters to detect clogging early. | No abnormal DP increase |  | Medium
AI-48 | Advanced Diagnostics and Predictive Monitoring | Check low flow alarm strategy | Confirm low flow transmitters or switches generate timely alarms. | Alarm strategy active |  | Medium
AI-49 | Advanced Diagnostics and Predictive Monitoring | Compare historical trends for drift | Review longer term data for gradual degradation patterns and maintenance planning. | No abnormal drift trend |  | High

## Summary
Analyzer Inspection Checklist Summary
Live counts and completion overview based on the checklist status column
Metric | Value | Priority | Count
Total items | 50 | Critical | 10
OK | 0 | High | 14
Not OK | 0 | Medium | 25
In Progress | 1 | Low | 0
Not Started | 0
Status and Priority Guide
Status | Meaning | Priority | Meaning
Not Started | Not yet checked | Low | Routine item
In Progress | Inspection underway | Medium | Important item
OK | Acceptable / passed | High | Operational impact
Not OK | Needs action or repair | Critical | Immediate attention

## Legend
Checklist Notes and Use Guide
Use the Status dropdown in column F and Priority dropdown in column G for each inspection point
Field | Purpose | Example | How to use
Status | Inspection result | OK / Not OK | Choose the result after checking each point
Priority | Operational importance | High / Critical | Use this to rank urgent items
Owner | Responsible person | Analyzer technician | Assign the person who will close the action
Remarks | Follow-up notes | Replace filter cartridge | Capture findings and actions
Acceptance Criteria | Pass standard | No leakage detected | Compare against the criteria column